放射線機器管理士測驗

Medical Imaging and Radiologic Systems Manager

2018年8月26日星期日

- Choose the item that can be measured by using test charts of the star-pattern camera method.

- mA Linearity

- mAs reciprocity

- kVp accuracy

- Focal spot test

- What is this test for?

- Slice thickness determination

- Uniformity determination

- Low contrast resolution measurements

- High contrast (spatial) resolution

- Choose the answer that correctly describes the discrepancy between the X-ray field and the light field.

- 0.5% of the source image distance (SID)

- 1.0% of the SID

- 2.0% of the SID

- 3.0% of the SID

- Choose the property that is not relevant to the performance of automatic exposure control.

- Response time

- Tube voltage

- Subject thickness

- Tube current

- The second half-value layer (HVL) of a photon beam is approximately the same as first HVL:

- For all X-ray tube generated photon beams

- Only if the energy is below 100 kVp

- Only if the beam is monoenergetic ( e.g., gamma rays)

- Never, it is always less

- List the following types of X-ray equipment in order of increasing focal spot size.

- CT, mammography, angiography, chest unit

- Mammography, angiography, chest unit, CT

- Angiography, CT, chest unit, angiography

- Chest unit, mammography, angiography, CT

- In X-ray tube, off-focus radiation?

- is low-energy bremsstrahlung radiation from the non-target part of the anode

- will not make up more than 1% of the tube output

- has no effect on image quality

- makes no contribution to patient dose

- A major limitation of many early full-field digital mammography units is the?

- Focal spot size

- Detector size

- Lack of compression capability

- Limited angulation

- In mammography, what would an SMPTE pattern be used for?

- Evaluation of system resolution

- Evaluation of focal spot size

- QC of laser image printer and display monitor

- View box QC

- The interventional reference point (defined by the International Electro technical Commission) is located?

- At 15 cm from the isocenter toward the image receptor

- At 15 cm from the isocenter toward the x-ray tube

- At isocenter

- At the port of the x-ray tube

- Which one of the following provides the best estimate of the maximum skin dose delivered during an interventional procedure?

- The displayed dose at the reference point

- The displayed kerma area product (KAP)

- The total number of DSA frames used for case

- The total fluoroscopic time used for case

- A 64-slices CT scanner can obtain sixty-four 0.625 mm wide slices in a single 0.5 second axial acquisition. How many 5 mm wide slices can it obtain in a single 1.0 second axial acquisition?

- 8

- 16

- 64

- 128

- In order to estimate the effective dose in mSv from a CT exam, which of the following is “not” needed?

- the length of the CT scan study

- a dose-length product to effective dose conversion factor for body part imaged

- the patient weight

- CTDIvol of the patient

- For MR imaging, the patient’s weight must be entered accurately during patient registration to?

- Prevent damage to the RF coils

- Prevent artifacts in the images

- Prevent damage to the patient table

- Enable calculation of the specific absorption rate (SAR)

- The width of an ultrasound beam measured perpendicular to the image plane determines the?

- Axial resolution

- Depth resolution

- Lateral resolution

- Slice thickness

- What is the ACR CT accreditation program reference dose for routine adult abdomen CT scan?

- 60 mGy

- 25 mGy

- 35mGy

- 10mGy

- In MRI, which gradient is on during the time the echo is measured?

- Slice selection

- Phase encoding

- Collimating

- Frequency encoding

- Which of the following is “not” a method of medical radioisotope production?

- Bombarding with neutrons in a reactor

- Bombarding with protons in a cyclotron

- Bombarding with neutrons in a cyclotron

- Elution of a metastable daughter from a parent

- Scintillation detector conversion efficiency is?

- The same as the detector detection efficiency

- The fraction of deposited energy that is converted into light

- The fraction of deposited energy that is converted into electrons

- Decrease with the scintillator atomic number

- The x-ray tube current is

- Determined by the anode temperature

- Adjusted through varying the filament current

- Usually space charge limited

- Usually 500 mA for mammography

- The resolving capability of the focal spot is best:

- Towards the anode side of the field

- Directly under the central axis (center of the field)

- Towards the cathode side of field

- Towards the left side of the field perpendicular to anode-cathode axis

- Geometric magnification can improve the detection of high contrast objects. The fundamental limitation on useful magnification is:

- Blurring due to focal spot size

- Blurring due to removal of the grid

- MTF of the image receptor

- H&D curve of the image receptor

- Low contrast detectability refers to the ability of system to distinguish:

- A calcified lung nodule

- A non-calcified lung nodule

- The size of large fracture

- The size of small fracture

- The image of a resolution test pattern placed in the center of the x-ray field and 4.5 cm from the image receptor shows a limiting resolution of 11 lp/mm. what resolution may be expected if the pattern is moved to the anode-side edge of the field?

- 1 lp/mm

- 10 lp/mm

- 20 lp/mm

- 12 lp/mm

- In a CT image, the Hounsfield number of a material:

- Is dependent on its linear attenuation coefficient

- Increase if the window width is increased

- Has units of per cm

- Increase with mA

- What is this phantom test for?

- Radiation dose

- Uniformity determination

- Low contrast resolution measurements

- High contrast (spatial) resolution

- A DR system specification sheet says the field of view is 43*43 cm, and the detector element size is 100*100 microns. What is the spatial resolution?

- 2.15 lp/mm

- 2.5 lp/mm

- 4.3 lp/mm

- 5.0 lp/mm

- What is this test for?

- CT number calibration

- Uniformity determination

- Low contrast resolution measurements

- High contrast (spatial) resolution

- In mammography, what would be the impact of changing a Mo/Mo to a Mo/Rh target/filter system (AGD is average glandular dose)

- The AGD would decrease and contrast would increase

- The AGD would decrease and contrast would decrease

- The AGD would increase and contrast would decrease

- The AGD would increase and contrast would increase

- What is the most common phosphor used in DR system?

- CsI

- NaI

- GdOS

- YGdO

- Which modality require the most shielding in the wall, floor and ceiling?

- DR unit

- FFDM mammography unit

- Fluoroscope unit

- CT scanner

- For the same collimation, at which pitch would helical mode CT have a higher patient dose than axial mode?

- 3

- 1.5

- 1.8

- 1

- A computed radiography image with 10 bits per pixel. How many shades of gray and byte/pixel will have?

- 256, 1

- 1024, 1

- 256, 2

- 1024, 2

- To keep patient dose as low as reasonably achievable for digital radiography exam using a CR or DR detector, the mAs is selected to provide acceptable:

- Image contrast

- Optical density

- Luminance on the image monitor

- Signal to noise ratio

- Choose the item that is least relevant to medical quality.

- Effectiveness

- Efficiency

- Accessibility

- Scientific and technical quality

- Choose the correct name of the stepwise approach consisting of setting an objective, making and executing a plan, and evaluating and managing the performance, so as to achieve the objective effectively; this method was developed in the 1930s and is still being used widely.

- CPDA

- DAPC

- PDCA

- ACPD

- Choose the correct range of total filtration for X-ray equipment with a nominal

maximum tube voltage of 150 kV or less.

- 1.5 mmAl or more

- 2.0 mmAl or more

- 2.5 mmAl or more

- 3.0 mmAl or more

- Choose the term that describes a way to provide image data at constant density regardless of the body shape of the patient by controlling exposure time and other factors.

- Generator

- Automatic exposure control (AEC)

- Half value layer (HVL)

- Timer

- Choose the correct combination of items that affect the quality of mammographic images.

a. Absorption difference in the breast   b. Size of focus  c. Exposure field  d. Filtration

- Only a

- a and b

- a, b and c

- All of a to d

- Choose the item that can be measured by using the phantom shown below.

- CT number calibration

- Uniformity determination

- Low contrast resolution measurements

- High contrast (spatial) resolution

- Choose the correct combination of items that are relevant to a multi-slice CT system.

a. RF coil  b. Radio shield  c. Cone beam  d. Data acquisition system (DAS)

e. Sensitivity time control (STC)

- a and b

- a and e

- b and c

- c and d

- Choose the answer that does not correctly describe the computed tomography dose index (CTDI) of a CT system.

- A pensile-type ion chamber is used.

- Phantoms of 16 and 32 cm in diameter are used.

- The Z-side dose profile of multi-slice scan is integrated.

- Data on patient exposure during the operation is provided.

- Choose the description appropriate for the image shown below.

- Spatial resolution measurement

- Low contrast resolution

- CT number measurement

- Dosimetry

- Choose the correct measure of the spatial resolution for a CT system.

- Winer spectrum (SD)

- cm-1/CT number

- CT number

- MTF (Lp/cm)

- Choose the item that does not cause artifacts in the MRI.

- Breathing motion

- Intestinal gas

- Cardiac motion

- Buttons on clothes

- Choose the item that should not be brought into an MR examination room.

- Oxygen cylinder

- Ceramic scissors

- Wheelchair made of non-magnetic substances

- Butterfly needle

- How many imitation lesions are there in the mammo phantom (it is made out of 50 % breast parenchymal/50 % fat and is pressed in to 4.5 cm)?

- Fiber 6 speck 5 mass 5

- Fiber 6 speck 6 mass 5

- Fiber5 speck 5 mass 6

- Fiber 5 speck 6 mass 6

- The MTF is one method to measure the:

- Low-contrast resolution

- High-contrast spatial resolution

- Attenuation

- Section Thickness

- Which of the following is not related with the performance of automatic exposure control device?

- The characteristic of response time

- The characteristic of kvp

- The characteristic of subject thickness

- The characteristic of mAs

- How often at least for Mammographic quality standards, phantom images with ACR accreditation phantom or equivalent phantom should be obtain?

- Daily.

- Weekly.

- Monthly.

- Yearly.

- Choose the item not related to computed radiography (CR).

- Cassette with intensifying screen

- Imaging plate

- Laser tube

- A/D converter

- Choose the item not related to a CT system.

- Single slice

- Multi-slice

- Quenching

- Helical scan

- Choose the correct statement about an X-ray CT value.

- Fat has a higher CT value than water

- It is determined from the reflectance of an X-ray beam

- It depends on the atomic number of the subject material

- It is a relative value, with a zero value for water

- Choose the item not relevant to the performance evaluation of a CT system.

- Spatial resolution

- Scintillator

- Contrast resolution

- Slice thickness

- Choose the item not relevant to the start-of-day checking of a CT system.

- Gantry operation and safety

- Couch operation and safety

- X-ray output dose measurement

- Safety mechanism operation

- Choose the item not relevant to the image noise of a CT system.

- Structural noise

- Acoustic noise

- Electric noise

- Quantization noise

- Choose the item not relevant to the dosimetry of a CT system.

- Low contrast resolution

- DLP

- MSAD

- CTDI

- Filtered back projection in CT refers to ?

- spatial resolution

- pre-patient collimation

- beam filtration

- image reconstruction

- Which type of test tool may be used to demonstrate pincushion distortion?

- Dosimeter

- Wire mesh tool

- Ion chamber

- Homogenous phantom

- Which of CT quality control is measured in the figure?

- Noise and Uniformity

- Spatial Resolution

- Slice Thickness

- Patient Dose

- Which of CT quality control is measured in the figure?

- Noise and Uniformity

- Radiation beam width

- Slice Thickness

- Patient Dose

- Image noise in CT scanning depends on?

- pixel size

- detector efficiency

- slice thickness

- all of the above

- The patient dose in CT is determined by?

- Pre-detector collimator

- Pre-patient collimator

- Post-patient collimator

- all of the above

- In MRI, signal to noise ratio can be increase by all of the following expect:

- switching from body to head coil

- increasing the number of acquisitions

- increasing the static magnetic field strength

- decreasing the slice thickness

- Chemical shift artifacts in MRI are:

- secondary to using gadolinium contrast

- seen in the frequency encoding direction

- not affected by the gradient strength

- more severe for lower magnetic field strength

- A 4 MHz ultrasound beam travels through 7.5 cm of soft tissue with an attenuation of 1 dB/cm/MHz. The original intensity is reduced by?

- 30 dB

- 100 dB

- 300 dB

- 3 dB

- If an ultrasound pulse is about three wavelengths long and wavelength in soft tissue is 0.33 mm. what is the achievable axial resolution in mm?

- 0.5

- 1.0

- 3.3

- 9.9

- The SI unit of absorbed dose from ionizing radiation is the?

- Becquerel

- MeV

- Gray

- Rem

- Sievert is a unit of?

- Exposure

- Equivalent dose

- Dose rate

- Energy

- The roentgen is a unit of:

- Radiation exposure

- Absorbed dose

- Equivalent dose

- Radioactivity

- If the linear attenuation coefficient is 0.5 cm-1, the HVL is ?

- 13.86 cm

- 1.386 cm

- 0.05 cm

- 0.0347 cm

- Which of following rays are electromagnetic waves?

- Beta

- Gamma

- Delta

- Alpha

- 1 Bq is equal to?

- 2.7*10-11Ci

- 3.7*1010Ci

- 3.7*10-10Ci

- 2.7*10-7Ci

- Which of the following types of radiation is not electromagnetic?

- Microwave

- Ultrasound

- Radiowaves

- Heat radiation

- An H&D curve illustrates the dependency of?

- Dose upon exposure

- Optical density upon dose

- Dose upon depth

- Film transmission upon depth

- Pixel size for full-field digital mammography units is ?

- 0.05 to 0.1 mm

- 0.2 to 0.5 mm

- 1 to 2 mm

- 0.8 to 1 mm

- In PACS, large-matrix display stations(for DR or CR) have all of the following except:

- A magnification function

- Interactive window and level function

- Ability to measure the linear attenuation coefficient accurately

- Ability to invert the gray scale value of the image

- To store images with 256shades of gray, each pixel in the image will require:

- 1 byte

- 6 bytes

- 1 bit

- 6 bits

- For a similar RF coil, 3T MRI always has which than 1.5T MRI?

- Higher resolution

- Higher SNR

- A shorter scan time

- Less artifacts

- In PACS, which of the following is a mass storage device

- DICOM

- NEMA

- RAID

- ISDN

- The effective energy of an x-ray beam?

- linearly increase with the atomic number

- affects subject contrast

- is equal to the kVp

- is proportional to the mAs

- What is this test for?

- Spatial resolution measurement

- Low contrast resolution

- Focal spot measurement

- Dosimetry

- What is this test for?

- Spatial resolution measurement

- Low contrast resolution

- Focal spot measurement

- Dosimetry

- What is this test for?

- Spatial resolution measurement

- Beam collimation

- Focal spot measurement

- Dosimetry

- What is the purpose of the directions in the image?

- Gray scale steps

- Contrast resolution

- Uniform background

- 5 % Contrast patches

- Choose the item relevant to CT resolution in the axial direction.

- Scan time

- Helical pitch

- Display pixel size

- Detector sensitivity

- Which of the following is correct standard range when using 250mAs in CT equipment QC?

- 10 %

- 2.15 %

- 20 %

- 25 %

- Which effect is true about inhomogeneity of CT number from characteristics of X-ray in an object in Computed Tomography?

- Photon effect

- Cupping effect

- Partial volume effect

- Beam hardening effect

- In phantom image exam of CT equipment, What is the standard range of CT number and Noise?

- 7 HU below, 7 HU below

- 10 HU below, 8 HU below

- 5 HU below, 10 HU below

- 7 HU below, 10 HU below

- Choose the correct action to be taken in advance of a power system maintenance outage in an area including an MRI examination room.

- In the case of a Resistive magnet MRI system, keeping the power to the Static magnetic field coil is not related with eliminating the magnetic field.

- In the case of a Permanent magnet MRI system, sudden change in temperature in the examination room is not related with magnetic field.

- In all types of MRI systems, because the Static magnet field disappears during an outage, there is not need to prohibit access to the MRI examination room.

- In the case of a Superconducting magnet MRI system, because helium consumption increases when the refrigerator stops, add a sufficient amount of helium to avoid the occurrence of quenching.